Protocol for Treating Stress Urinary Incontinence (SUI) Using Low-Intensity

Shockwaves

Device: MoreNovaPro - Li50F Linear Transducer

Document Version: 2.0

Date: 16/12/2025

1. Purpose and Scope

This protocol provides a standardized treatment plan for using low-intensity shockwaves to

manage stress urinary incontinence (SUI). It is intended for healthcare providers using the

MoreNova Device with Li50F Transducer to target the tissues surrounding the urethra and

sphincter muscles. This protocol applies to women aged 20-75 diagnosed with SUI.

2. Introduction

Low-Intensity Extracorporeal Shockwave Therapy (LiESWT) is a non-invasive technology

that uses acoustic waves to stimulate tissue regeneration, enhance blood circulation, and

strengthen muscle function. It is clinically proven for treating stress urinary incontinence

(SUI) by targeting weakened pelvic floor muscles and the urethral sphincter. For further

information on clinical studies and efficacy, visit our Clinical Data Center.

Device Overview:

MoreNova is a portable Electromagnetic (EM) Low-Intensity Shockwave (LISW) device

designed to reduce inflammation, dissolve calcifications, and stimulate angiogenesis,

improving outcomes for various conditions in men and women.

3. Background

Condition Overview:

•

•

SUI is an involuntary leakage of urine due to increased abdominal pressure during

activities like coughing, sneezing, or exercising.

It often results from weakened pelvic floor muscles and urethral sphincter

dysfunction.

Mechanism of Action:

•

Low-intensity shockwave therapy stimulates neovascularization, enhances blood

flow, and promotes tissue regeneration, improving pelvic floor strength and

urethral function.

4. Patient Selection Criteria

Inclusion Criteria:

•

•

•

Women aged 20-75 years.

Diagnosed with SUI by a gynecologist for at least 3 months.

Persistent SUI symptoms despite conservative management.

Exclusion Criteria:

•

•

•

•

•

•

•

•

History of surgery in the vulvar or vestibular region.

Blood coagulation disorders.

Active pelvic inflammatory conditions.

Vaginal infections (e.g., yeast infection or bacterial vaginosis).

Pregnancy or lactation.

History of pelvic or external genitalia irradiation within the past year.

Chronic neurological disorders cause localized pain.

Presence of an electromagnetic-sensitive medical device.

5. Pre-Treatment Screening

To determine the patient's suitability for LiESWT, conduct the following:

•

•

•

Comprehensive medical history review.

Physical examination to assess pelvic floor condition.

Diagnostic tests: Urinalysis to rule out infections, and pelvic ultrasound for

anatomical assessment, if indicated.

6. Treatment Protocol

Preparation:

•

•

Conduct a thorough medical history and gynecological examination.

Counsel the patient about the procedure, benefits, risks, and expected outcomes.

Procedure:

•

•

•

•

Number of Sessions: 6-8 sessions based on severity.

Session Frequency: Twice per week.

Session Duration: 20 minutes. (angled rotation beneath the urethra)

Applications Per Session: 5.

The rotation pattern:

1

2

3

4

5

•

.

.

.

.

.

-30°

-15°

0° (beneath the urethra)

15°

30°

Shocks Per Application: 500.

•

Total Shocks Per Session: 2500.

Treatment Parameters:

•

•

Energy (HV) Settings: 2.0 kV.

Frequency (FRQ): 2.0 Hz.

Post-Treatment Care:

•

•

Advise patients to avoid strenuous physical activity for 24 hours.

Recommend hydration and report any adverse events.

7. Treatment Areas and Techniques

The treatment focuses on the following anatomical areas:

•

•

Pelvic floor muscles.

Urethral sphincter.

Ensure correct Transducer placement beneath the urethra and apply gentle pressure. Rotate

the Transducer in 15° increments to cover all necessary areas, following the predefined

sequence.

8. Patient Comfort and Safety

To ensure patient comfort and safety during the procedure:

•

•

•

Position the patient in a supine or lithotomy position for optimal access.

Use conductive gel for effective shockwave transmission and minimal discomfort.

Continuously monitor patient feedback during the session and adjust settings if

needed.

9. Treatment Outcomes and Follow-Up

Expected Results:

Patients typically experience symptom improvement within 3–6 weeks of starting

treatment. Regular assessments are recommended to evaluate progress.

Follow-Up Visits:

Schedule bi-weekly follow-ups to monitor patient outcomes and adjust the treatment plan

as needed.

Long-Term Maintenance:

For sustained improvement, consider maintenance sessions every 3–6 months based on

individual needs.

10. Visuals and Diagrams

Li50F

•

Anatomical areas being treated.

•

Rotation pattern for Transducer placement.

11. Quality Control

Device Calibration:

•

•

Verify device settings before each session.

The MoreNova Device recognizes preprogrammed parameters, ensuring correct

settings are automatically applied.

Clinician Training:

•

•

Ensure all operators are thoroughly trained in the use of the MoreNova Device and

Li50F Transducer.

Clinicians should be familiar with this protocol and review the "Instructions for Use"

provided by the manufacturer.

Checklist for Each Session:

•

•

•

Confirm that the patient meets the inclusion criteria and has no contraindications.

Ensure the device and Transducer are clean, sanitized, and function correctly.

Confirm that the Transducer placement aligns with the anatomical areas

specified in the protocol.

12. References

For clinical studies and additional resources, visit the Initia Clinical Data Center.